## Calculations
Re-entrant cylindrical cavity caculations. Single: one cylindrical post with gap at one end. Double: two posts of equal length at each end with gap in middle. Tracy Hall Jr, Hall Labs, October 2022
Use goal seek to find the gap d, in first column, that gives f0 = 100 MHz
|  |  |  | Inside radius of cavity | radius of post | inside height of cavity | gap for 100 MHz | rPost/ rCav | full range gap chart x | full range freq chart left | full range Q chart right | low range gap chart x | low range freq chart left | low range Q chart right |  |  | geo metric gap factor
|  |  |  | r2 | r1 | l | d
Input parameters |  | ` |  | 0.025 | 0.01 | 0.075 | 1.5223e-05 | [9 empty] | 1.5644
Saved inputs |  |  |  | 0.025 | 0.01 | 0.075 | 1.5223e-05 | 0.4 | 0.08 | 5500 | 11000 | 0.0009 | 1000 | 5000
|  |  |  | 0.025 | 0.015 | 0.075 | 1.9197e-05 | 0.6 | 0.075 | 4500 | 4500 | 0.001 | 700 | 3500
|  |  |  | 0.05 | 0.03 | 0.15 | 0.00015877 | 0.6 | 0.15 | 2400 | 1200 | 0.006 | 600 | 6000
Summary of Results | [6 empty] | (r1/r2) = | 0.4 | [11 empty] | max d
d | 1.5223e-05 | 2.3815e-05 | 3.7256e-05 | 5.8283e-05 | 9.1177e-05 | 0.00014264 | 0.00022314 | 0.00034908 | 0.00054609 | 0.0008543 | 0.0013365 | 0.0020907 | 0.0032707 | 0.0051167 | 0.0080045 | 0.012522 | 0.01959 | 0.030646 | 0.047942 | 0.075
f0 single, MHz | 99.5811 | 124.0464 | 154.2354 | 191.2663 | 236.317 | 290.5171 | 354.7806 | 429.5977 | 514.8508 | 609.7739 | 713.1851 | 824.0521 | 942.3158 | 1069.8414 | 1211.5382 | 1377.1748 | 1585.5548 | 1876.5271 | 2355.6058 | 3473.9432
f0 double, MHz | 100 | 124.8171 | 155.6433 | 193.8146 | 240.8739 | 298.5379 | 368.6137 | 452.8476 | 552.7086 | 669.1491 | 802.457 | 952.3765 | 1118.6867 | 1302.3473 | 1507.2728 | 1743.0481 | 2030.0337 | 2412.1789 | 3000.2074 | 4198.1474
Q single | 1682.9512 | 1878.3444 | 2094.4749 | 2332.3964 | 2592.5709 | 2874.5444 | 3176.604 | 3495.5385 | 3826.6951 | 4164.5451 | 4503.8536 | 4841.2831 | 5177.0384 | 5516.2361 | 5870.1839 | 6258.6069 | 6715.4292 | 7305.6787 | 8185.2972 | 9940.1858
Q double | 1686.4872 | 1884.1706 | 2104.013 | 2347.8827 | 2617.4475 | 2913.9552 | 3237.9405 | 3588.882 | 3964.891 | 4362.5921 | 4777.4266 | 5204.6015 | 5640.7568 | 6086.2047 | 6547.5537 | 7041.0566 | 7598.6207 | 8283.0045 | 9237.5921 | 10927.2778
single λ , m | 3.0105 | 2.4168 | 1.9437 | 1.5674 | 1.2686 | 1.0319 | 0.84501 | 0.69784 | 0.58229 | 0.49165 | 0.42036 | 0.3638 | 0.31814 | 0.28022 | 0.24745 | 0.21769 | 0.18908 | 0.15976 | 0.12727 | 0.086297
dble λ , m | 2.9979 | 2.4019 | 1.9262 | 1.5468 | 1.2446 | 1.0042 | 0.8133 | 0.66202 | 0.54241 | 0.44802 | 0.37359 | 0.31478 | 0.26799 | 0.23019 | 0.1989 | 0.17199 | 0.14768 | 0.12428 | 0.099924 | 0.071411
sngl MHz/µm | 2.8475 | 2.246 | 1.7611 | 1.3696 | 1.0533 | 0.79828 | 0.59408 | 0.43272 | 0.30798 | 0.21448 | 0.14698 | 0.10022 | 0.069083 | 0.049067 | 0.036664 | 0.029485 | 0.026318 | 0.027699 | 0.041331
dbl MHz/µm | 2.8885 | 2.2935 | 1.8154 | 1.4306 | 1.1206 | 0.87048 | 0.66886 | 0.50687 | 0.3778 | 0.27648 | 0.19876 | 0.14094 | 0.099493 | 0.070962 | 0.052189 | 0.040607 | 0.034564 | 0.033998 | 0.044273
`
https://en.wikipedia.org/wiki/Electrical_resistivity_and_conductivity#Resistivity_and_conductivity_of_various_materials
σ, Conductivity, Siemens/meter |  |  |  | Ag | 61728000 |  | Cu | 58005000 |  | Au | 40984000 |  | Al | 35386000 |  | yellow brass | 15659000
Single: Calculations for a re-entrant cylindrical cavity with cylindrical post or truncated conical post.
Source: J J Barroso et al, "Reentrant Klystron Cavity as an electromechanical transducer," 2004. | [9 empty] | http://mtc-m16.sid.inpe.br/col/sid.inpe.br/marciana/2004/01.16.14.41/doc/JJBarroso_Proc_IMOC2003.pdf
Based upon K. Fujisawa, "General Treatment of Klystron Resonant Cavities," IRE Transactions on Microwave Theory & Techniques, Oct. 1958, pp 344-358 | [15 empty] | https://ur.booksc.me/book/18439819/414edd
Permeability of free space µ0
[15 empty] | 1.2566e-06 | Henry/meter
Permittivity of free space ε0
[15 empty] | 8.8542e-12 | Farad/meter
Speed of light c
[15 empty] | 299792458 | meter/sec
Deuterium hyperfine line
[15 empty] | 327.3844 | MHz
Obtain minimum d by goal seek so f0=100 MHz. Geometrically progress d by dmult so dmax = length
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |  | 14 | 15 | 16 | 17 | 18 | 19 | 20
r2 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025
r1 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
r1/r2 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
l | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075
d | 1.5223e-05 | 2.3815e-05 | 3.7256e-05 | 5.8283e-05 | 9.1177e-05 | 0.00014264 | 0.00022314 | 0.00034908 | 0.00054609 | 0.0008543 | 0.0013365 | 0.0020907 | 0.0032707 | 0.0051167 | 0.0080045 | 0.012522 | 0.01959 | 0.030646 | 0.047942 | 0.075
α, deg | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90
α, rad | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708
͞͞r0 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
r0 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
lM | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243 | 0.038243
C1, Farad | 3.1269e-12 | 2.9684e-12 | 2.8099e-12 | 2.6514e-12 | 2.493e-12 | 2.3345e-12 | 2.176e-12 | 2.0175e-12 | 1.859e-12 | 1.7005e-12 | 1.542e-12 | 1.3835e-12 | 1.225e-12 | 1.0666e-12 | 9.0807e-13 | 7.4958e-13 | 5.9109e-13 | 4.326e-13 | 2.7411e-13 | 1.1562e-13
C0, Farad | 1.8272e-10 | 1.168e-10 | 7.4663e-11 | 4.7726e-11 | 3.0508e-11 | 1.9501e-11 | 1.2466e-11 | 7.9685e-12 | 5.0937e-12 | 3.256e-12 | 2.0813e-12 | 1.3304e-12 | 8.5046e-13 | 5.4364e-13 | 3.4751e-13 | 2.2214e-13 | 1.4199e-13 | 9.0767e-14 | 5.8021e-14 | 3.7088e-14
L, Henry | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08
f0, MHz | 99.5811 | 124.0464 | 154.2354 | 191.2663 | 236.317 | 290.5171 | 354.7806 | 429.5977 | 514.8508 | 609.7739 | 713.1851 | 824.0521 | 942.3158 | 1069.8414 | 1211.5382 | 1377.1748 | 1585.5548 | 1876.5271 | 2355.6058 | 3473.9432
λ, meter | 3.0105 | 2.4168 | 1.9437 | 1.5674 | 1.2686 | 1.0319 | 0.84501 | 0.69784 | 0.58229 | 0.49165 | 0.42036 | 0.3638 | 0.31814 | 0.28022 | 0.24745 | 0.21769 | 0.18908 | 0.15976 | 0.12727 | 0.086297
Cu Skin δ, m | 6.6222e-06 | 5.9333e-06 | 5.321e-06 | 4.7782e-06 | 4.2987e-06 | 3.8771e-06 | 3.5084e-06 | 3.1883e-06 | 2.9124e-06 | 2.6761e-06 | 2.4745e-06 | 2.302e-06 | 2.1527e-06 | 2.0204e-06 | 1.8985e-06 | 1.7807e-06 | 1.6596e-06 | 1.5255e-06 | 1.3616e-06 | 1.1212e-06
R(surface), Ω | 0.0026034 | 0.0029056 | 0.00324 | 0.003608 | 0.0040105 | 0.0044467 | 0.0049139 | 0.0054073 | 0.0059195 | 0.0064422 | 0.006967 | 0.007489 | 0.0080084 | 0.0085331 | 0.0090806 | 0.0096815 | 0.010388 | 0.011301 | 0.012662 | 0.015377
Geom Factor | 4.3813 | 5.4578 | 6.786 | 8.4153 | 10.3974 | 12.7821 | 15.6095 | 18.9013 | 22.6523 | 26.8287 | 31.3785 | 36.2564 | 41.4598 | 47.0706 | 53.3049 | 60.5926 | 69.7608 | 82.5629 | 103.6413 | 152.8456
Q | 1682.9512 | 1878.3444 | 2094.4749 | 2332.3964 | 2592.5709 | 2874.5444 | 3176.604 | 3495.5385 | 3826.6951 | 4164.5451 | 4503.8536 | 4841.2831 | 5177.0384 | 5516.2361 | 5870.1839 | 6258.6069 | 6715.4292 | 7305.6787 | 8185.2972 | 9940.1858
Geometric Factor and Q for single re-entrant cavity, practical smallest gap 3 µm: | [7 empty] | 2013 Le Floch et al_LCR & finite element calc tunable (2-22 GHz) re-entrant microwave cavity
https://arxiv.org/ftp/arxiv/papers/1308/1308.2755.pdf
Skin effect formula; 4 skin depths carries 98% of the current: | [7 empty] | https://en.wikipedia.org/wiki/Skin_effect#Cause
For constant gap and any length, minimum frequency is found when r1/r2 ~0.6
If the length is decreased, decreasing the gap by the same fraction keeps the frequency approximately the same (for l/2; 50 becomes 50.03, 3000 becomes 3610)
Double: Calculations for a symmetric, doubly re-entrant cylindrical cavity with gap centered in cavity
"Effect of Beam-tunnels on Resonant Frequency of Cylindrical Reentrant Cavity," Baniswal et al, Defence Science Journal, Vol. 71, No. 3, May 2021, pp. 332-336
https://publications.drdo.gov.in/ojs/index.php/dsj/article/view/16814/7540
set N=1, rb=0, then
Cg = ε0 * π * r02 / d
(Use same values as for single re-entrant cavity)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
r1 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025
r0/r1 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
r0 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
h | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075
d | 1.5223e-05 | 2.3815e-05 | 3.7256e-05 | 5.8283e-05 | 9.1177e-05 | 0.00014264 | 0.00022314 | 0.00034908 | 0.00054609 | 0.0008543 | 0.0013365 | 0.0020907 | 0.0032707 | 0.0051167 | 0.0080045 | 0.012522 | 0.01959 | 0.030646 | 0.047942 | 0.075
Cr, Farad | 1.5731e-12 | 1.4939e-12 | 1.4146e-12 | 1.3354e-12 | 1.2561e-12 | 1.1769e-12 | 1.0977e-12 | 1.0184e-12 | 9.3917e-13 | 8.5992e-13 | 7.8068e-13 | 7.0143e-13 | 6.2219e-13 | 5.4295e-13 | 4.637e-13 | 3.8446e-13 | 3.0521e-13 | 2.2597e-13 | 1.4672e-13 | 6.748e-14
Cg, Farad | 1.8272e-10 | 1.168e-10 | 7.4663e-11 | 4.7726e-11 | 3.0508e-11 | 1.9501e-11 | 1.2466e-11 | 7.9685e-12 | 5.0937e-12 | 3.256e-12 | 2.0813e-12 | 1.3304e-12 | 8.5046e-13 | 5.4364e-13 | 3.4751e-13 | 2.2214e-13 | 1.4199e-13 | 9.0767e-14 | 5.8021e-14 | 3.7088e-14
L, Henry | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08 | 1.3744e-08
f0, MHz | 100 | 124.8171 | 155.6433 | 193.8146 | 240.8739 | 298.5379 | 368.6137 | 452.8476 | 552.7086 | 669.1491 | 802.457 | 952.3765 | 1118.6867 | 1302.3473 | 1507.2728 | 1743.0481 | 2030.0337 | 2412.1789 | 3000.2074 | 4198.1474
λ, meter | 2.9979 | 2.4019 | 1.9262 | 1.5468 | 1.2446 | 1.0042 | 0.8133 | 0.66202 | 0.54241 | 0.44802 | 0.37359 | 0.31478 | 0.26799 | 0.23019 | 0.1989 | 0.17199 | 0.14768 | 0.12428 | 0.099924 | 0.071411
Cu Skin δ, m | 6.6083e-06 | 5.9149e-06 | 5.2969e-06 | 4.7467e-06 | 4.2579e-06 | 3.8246e-06 | 3.4419e-06 | 3.1054e-06 | 2.8109e-06 | 2.5546e-06 | 2.3328e-06 | 2.1413e-06 | 1.9758e-06 | 1.8312e-06 | 1.7021e-06 | 1.5828e-06 | 1.4667e-06 | 1.3455e-06 | 1.2065e-06 | 1.0199e-06
R(surface), Ω | 0.0026088 | 0.0029146 | 0.0032547 | 0.003632 | 0.0040489 | 0.0045076 | 0.0050088 | 0.0055517 | 0.0061333 | 0.0067485 | 0.0073902 | 0.008051 | 0.0087257 | 0.0094148 | 0.010128 | 0.010892 | 0.011754 | 0.012813 | 0.01429 | 0.016903
Geom Factor | 4.3998 | 5.4917 | 6.848 | 8.5274 | 10.5979 | 13.135 | 16.2182 | 19.9243 | 24.3179 | 29.441 | 35.3063 | 41.9024 | 49.2197 | 57.3003 | 66.3166 | 76.6902 | 89.3169 | 106.1304 | 132.0023 | 184.709
Q | 1686.4872 | 1884.1706 | 2104.013 | 2347.8827 | 2617.4475 | 2913.9552 | 3237.9405 | 3588.882 | 3964.891 | 4362.5921 | 4777.4266 | 5204.6015 | 5640.7568 | 6086.2047 | 6547.5537 | 7041.0566 | 7598.6207 | 8283.0045 | 9237.5921 | 10927.2778
Check formulas for single reentrant cavity
| Compare f0, this calc, with Fujisawa's measurements, cylindrical posts, Fig 13, p 354 | [8 empty] | Compare f0, this calc, with Fujisawa's measurements, truncated cones, Fig 15, p 355
Example | 61 | 61 | 61 | 72 | 72 | 72 | 79 | 79 | 79 | 91 | 91 | 91 | 97 | 97 | 97 | 101 | 101 | 101
r2 | 0.008925 | 0.008925 | 0.008925 | 0.014 | 0.014 | 0.014 | 0.01995 | 0.01995 | 0.01995 | 0.014 | 0.014 | 0.014 | 0.01995 | 0.01995 | 0.01995 | 0.01995 | 0.01995 | 0.01995
l | 0.0099 | 0.0099 | 0.0099 | 0.021 | 0.021 | 0.021 | 0.0209 | 0.0209 | 0.0209 | 0.006063 | 0.006188 | 0.00633 | 0.011517 | 0.012567 | 0.014057 | 0.015335 | 0.017445 | 0.019956
r1 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.010525 | 0.010525 | 0.010525 | 0.01035 | 0.01035 | 0.01035 | 0.008205 | 0.008205 | 0.008205
d | 0.000716 | 0.000845 | 0.000991 | 0.007294 | 0.008169 | 0.008938 | 0.0097 | 0.01069 | 0.011615 | 0.000563 | 0.000688 | 0.00083 | 0.003217 | 0.004267 | 0.005757 | 0.005235 | 0.007345 | 0.009856
α, deg | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 47.58 | 47.58 | 47.58 | 59.7 | 59.7 | 59.7 | 75.01 | 75.01 | 75.01
α, rad | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 0.83043 | 0.83043 | 0.83043 | 1.042 | 1.042 | 1.042 | 1.3092 | 1.3092 | 1.3092
͞͞r0 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0054993 | 0.0054993 | 0.0054993 | 0.0054999 | 0.0054999 | 0.0054999 | 0.0055006 | 0.0055006 | 0.0055006 |  | 1
r0 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0055 | 0.0049848 | 0.0048706 | 0.0047409 | 0.00362 | 0.0030064 | 0.0021358 | 0.0040989 | 0.0035339 | 0.0028615
lM | 0.0052379 | 0.0052379 | 0.0052379 | 0.011328 | 0.011328 | 0.011328 | 0.012704 | 0.012704 | 0.012704 | 0.006247 | 0.0063389 | 0.0064431 | 0.011036 | 0.011633 | 0.012478 | 0.011504 | 0.012434 | 0.013551
C1, Farad | 5.8243e-13 | 5.5016e-13 | 5.1911e-13 | 2.8054e-13 | 2.5847e-13 | 2.4094e-13 | 2.4735e-13 | 2.2842e-13 | 2.1226e-13 | 8.5837e-13 | 7.9246e-13 | 7.3294e-13 | 4.4758e-13 | 3.9817e-13 | 3.5942e-13 | 3.3758e-13 | 2.8984e-13 | 2.5836e-13 |  | `
C0, Farad | 1.1752e-12 | 9.9579e-13 | 8.4908e-13 | 1.1536e-13 | 1.03e-13 | 9.4142e-14 | 8.6747e-14 | 7.8713e-14 | 7.2444e-14 | 1.4942e-12 | 1.2227e-12 | 1.0135e-12 | 2.6155e-13 | 1.9719e-13 | 1.4615e-13 | 1.6077e-13 | 1.1458e-13 | 8.5392e-14
L, Henry | 9.5853e-10 | 9.5853e-10 | 9.5853e-10 | 3.9241e-09 | 3.9241e-09 | 3.9241e-09 | 5.3859e-09 | 5.3859e-09 | 5.3859e-09 | 3.4596e-10 | 3.5309e-10 | 3.6119e-10 | 1.5116e-09 | 1.6494e-09 | 1.845e-09 | 2.725e-09 | 3.0999e-09 | 3.5461e-09
f0, GHz | 3.8775 | 4.1345 | 4.3948 | 4.0379 | 4.2258 | 4.3891 | 3.7519 | 3.9132 | 4.0644 | 5.5788 | 5.9665 | 6.3368 | 4.8612 | 5.0789 | 5.2112 | 4.3189 | 4.4949 | 4.5585
Fujisawa f0 | 3.8 | 4 | 4.2 | 3.8 | 4 | 4.2 | 3.8 | 4 | 4.2 | 3.8 | 4 | 4.2 | 3.8 | 4 | 4.2 | 3.8 | 4 | 4.2
λ, meter | 0.077316 | 0.072511 | 0.068215 | 0.074244 | 0.070943 | 0.068304 | 0.079903 | 0.076611 | 0.07376 | 0.053738 | 0.050246 | 0.04731 | 0.061671 | 0.059027 | 0.057529 | 0.069415 | 0.066696 | 0.065766
Note: Fujisawa's example 90 is ~3 times too high, and examples 95 & 96 cause div/0, although all have consistent geometry, with rbar0 = .0055m
Calculations with a right cylindrical post are in much better agreement with experiment than those for the truncated cylinder. A right cylindrical post is also easier to work with.
Check against 2013 Le Floch et al_LCR & finite element calc tunable (2-22 GHz) re-entrant microwave cavity: Rcav =4.97 mm, Rpost=0.5mm, height =1.4mm, σ Cu = 3.3E7 Seimens/m
gap |  |  | 5.058e-06 | 7.311e-06 | 1e-05 |  |  |  |  | https://arxiv.org/ftp/arxiv/papers/1308/1308.2755.pdf
f (LCR), GHz |  |  | 5.113 | 6.049 | 6.949
Q (LCR) |  |  | 686.2 | 747.5 | 802.7 |  |  |  |  | The Quality Factor Q is the number of cycles, in radians/sec
|  |  | OK | OK | OK |  |  |  |  | that persist in the cavity until the stored energy drops to 1/e = | [6 empty] | 36.79%
r2 = Rcav | 0.00497 | 0.00497 | 0.00497 | 0.00497 | 0.00497 | 0.00497 | 0.00497 | 0.00497
r1 = Rpost | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 |  | Geometric Factor GF
r1/r2 | 0.1006 | 0.1006 | 0.1006 | 0.1006 | 0.1006 | 0.1006 | 0.1006 | 0.1006
l = height | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014 | 0.0014
d | 7.1788e-07 | 2.9894e-06 | 5.058e-06 | 7.311e-06 | 1e-05 | 2.359e-05 | 7.2559e-05 | 0.00027208
α, deg | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90
α, rad | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708
͞͞r0 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005
r0 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005
lM | 0.0023421 | 0.0023421 | 0.0023421 | 0.0023421 | 0.0023421 | 0.0023421 | 0.0023421 | 0.0023421 |  | Q = GF / Rsurface
C1, Farad | 1.6097e-13 | 1.3571e-13 | 1.264e-13 | 1.1988e-13 | 1.1433e-13 | 9.9131e-14 | 7.9234e-14 | 5.5829e-14
C0, Farad | 9.6869e-12 | 2.3263e-12 | 1.3749e-12 | 9.5118e-13 | 6.9541e-13 | 2.9478e-13 | 9.584e-14 | 2.5559e-14
L, Henry | 6.4304e-10 | 6.4304e-10 | 6.4304e-10 | 6.4304e-10 | 6.4304e-10 | 6.4304e-10 | 6.4304e-10 | 6.4304e-10
f0, MHz | 2000 | 4000 | 5122.3958 | 6064.5144 | 6974.773 | 10000 | 14999.9992 | 21999.9995
λ, meter | 0.1499 | 0.074948 | 0.058526 | 0.049434 | 0.042982 | 0.029979 | 0.019986 | 0.013627
Geometric Fctr | 6.6154 | 13.2309 | 16.9434 | 20.0597 | 23.0706 | 33.0772 | 49.6158 | 72.7698
skin depth δ | 1.9591e-06 | 1.3853e-06 | 1.2241e-06 | 1.125e-06 | 1.0491e-06 | 8.7612e-07 | 7.1535e-07 | 5.9068e-07
R surface, Ω | 0.015468 | 0.021875 | 0.024755 | 0.026935 | 0.028886 | 0.034588 | 0.042361 | 0.051302
Q | 427.6816 | 604.8331 | 684.4506 | 744.7381 | 798.6759 | 956.3251 | 1171.2542 | 1418.4593
d | 7.1788e-07 | 2.9894e-06 | 5.058e-06 | 7.311e-06 | 1e-05 | 2.359e-05 | 7.2559e-05 | 0.00027208
f0, MHz | 2000 | 4000 | 5122.3958 | 6064.5144 | 6974.773 | 10000 | 14999.9992 | 21999.9995
Q | 427.6816 | 604.8331 | 684.4506 | 744.7381 | 798.6759 | 956.3251 | 1171.2542 | 1418.4593

## rCav .025
Re-entrant cylindrical RESULTS (no formulas). Single: one cylindrical post with gap at one end. Double: two posts of equal length at each end with gap in middle.
|  |  |  | Inside radius of cavity | radius of post | inside height of cavity | gap for 100 MHz | gap geo-metric spacer | full range chart x | full range chart left | full range chart right | low range chart x | low range chart left | low range chart right
|  |  |  | r2 | r1 | l | d | dmult
Input parameters |  |  |  | 0.025 | 0.01 | 0.075 | 0.00015877 | 1.4342 | 0.075 | 4500 | 9000 | 0.001 | 700 | 3500
Summary of Results | [6 empty] | (r1/r2) = | 0.6 | [11 empty] | max d
d | 1.9197e-05 | 2.9667e-05 | 4.5848e-05 | 7.0854e-05 | 0.0001095 | 0.00016922 | 0.00026151 | 0.00040413 | 0.00062455 | 0.00096518 | 0.0014916 | 0.0023051 | 0.0035623 | 0.0055052 | 0.0085077 | 0.013148 | 0.020319 | 0.0314 | 0.048526 | 0.074992
f0 single | 100 | 123.9155 | 153.3324 | 189.3533 | 233.1877 | 286.0849 | 349.2193 | 423.5299 | 509.5431 | 607.2553 | 716.1915 | 835.7511 | 965.8937 | 1108.1374 | 1266.9064 | 1451.6465 | 1681.2642 | 1996.2486 | 2502.2637 | 3627.6101
f0 double | 100.3457 | 124.5413 | 154.4584 | 191.3637 | 236.7414 | 292.2864 | 359.8659 | 441.4332 | 538.8881 | 653.8991 | 787.7479 | 941.3228 | 1115.4446 | 1311.7316 | 1534.2287 | 1792.2321 | 2105.713 | 2518.4127 | 3140.144 | 4357.3065
Q single | 1285.2586 | 1430.7155 | 1591.5028 | 1768.5891 | 1962.6521 | 2173.8918 | 2401.8153 | 2645.0418 | 2901.2218 | 3167.2048 | 3439.5772 | 3715.6001 | 3994.433 | 4278.4582 | 4574.7013 | 4896.8964 | 5269.9739 | 5742.4575 | 6429.2012 | 7741.0663
Q double | 1287.4783 | 1434.3239 | 1597.336 | 1777.953 | 1977.5503 | 2197.3275 | 2438.1523 | 2700.368 | 2983.5943 | 3286.5924 | 3607.3153 | 3943.2996 | 4292.5413 | 4654.928 | 5034.2587 | 5441.1112 | 5897.8005 | 6449.9142 | 7202.2015 | 8483.9781
single λ , m | 2.9979 | 2.4193 | 1.9552 | 1.5832 | 1.2856 | 1.0479 | 0.85846 | 0.70784 | 0.58836 | 0.49368 | 0.41859 | 0.35871 | 0.31038 | 0.27054 | 0.23663 | 0.20652 | 0.17831 | 0.15018 | 0.11981 | 0.082642
dble λ , m | 2.9876 | 2.4072 | 1.9409 | 1.5666 | 1.2663 | 1.0257 | 0.83307 | 0.67913 | 0.55632 | 0.45847 | 0.38057 | 0.31848 | 0.26876 | 0.22855 | 0.1954 | 0.16727 | 0.14237 | 0.11904 | 0.095471 | 0.068802
sngl MHz/µm | 2.2842 | 1.818 | 1.4405 | 1.1343 | 0.88576 | 0.68408 | 0.52102 | 0.39023 | 0.28686 | 0.20694 | 0.14697 | 0.10352 | 0.073213 | 0.052879 | 0.039814 | 0.032021 | 0.028424 | 0.029547 | 0.04252
dbl MHz/µm | 2.3109 | 1.8489 | 1.4759 | 1.1743 | 0.9301 | 0.73224 | 0.57189 | 0.44214 | 0.33764 | 0.25427 | 0.18878 | 0.1385 | 0.10103 | 0.074103 | 0.055603 | 0.043716 | 0.037241 | 0.036304 | 0.045989
SINGLE post
r2 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025
r1 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015
r1/r2 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
l | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075
d | 1.9197e-05 | 2.9667e-05 | 4.5848e-05 | 7.0854e-05 | 0.0001095 | 0.00016922 | 0.00026151 | 0.00040413 | 0.00062455 | 0.00096518 | 0.0014916 | 0.0023051 | 0.0035623 | 0.0055052 | 0.0085077 | 0.013148 | 0.020319 | 0.0314 | 0.048526 | 0.074992
α, deg | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90
α, rad | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708
͞͞r0 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015
r0 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015
lM | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832 | 0.037832
C1, Farad | 4.5614e-12 | 4.3302e-12 | 4.0989e-12 | 3.8677e-12 | 3.6364e-12 | 3.4052e-12 | 3.1739e-12 | 2.9427e-12 | 2.7114e-12 | 2.4802e-12 | 2.249e-12 | 2.0177e-12 | 1.7865e-12 | 1.5552e-12 | 1.324e-12 | 1.0927e-12 | 8.6149e-13 | 6.3024e-13 | 3.99e-13 | 1.6775e-13
C0, Farad | 3.2602e-10 | 2.1096e-10 | 1.3651e-10 | 8.8332e-11 | 5.7158e-11 | 3.6986e-11 | 2.3933e-11 | 1.5487e-11 | 1.0021e-11 | 6.4845e-12 | 4.196e-12 | 2.7151e-12 | 1.7569e-12 | 1.1369e-12 | 7.3565e-13 | 4.7602e-13 | 3.0803e-13 | 1.9932e-13 | 1.2898e-13 | 8.3457e-14
L, Henry | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09
f0, MHz | 100 | 123.9155 | 153.3324 | 189.3533 | 233.1877 | 286.0849 | 349.2193 | 423.5299 | 509.5431 | 607.2553 | 716.1915 | 835.7511 | 965.8937 | 1108.1374 | 1266.9064 | 1451.6465 | 1681.2642 | 1996.2486 | 2502.2637 | 3627.6101
λ, meter | 2.9979 | 2.4193 | 1.9552 | 1.5832 | 1.2856 | 1.0479 | 0.85846 | 0.70784 | 0.58836 | 0.49368 | 0.41859 | 0.35871 | 0.31038 | 0.27054 | 0.23663 | 0.20652 | 0.17831 | 0.15018 | 0.11981 | 0.082642
Cu Skin δ, m | 6.6083e-06 | 5.9364e-06 | 5.3367e-06 | 4.8023e-06 | 4.3275e-06 | 3.907e-06 | 3.5362e-06 | 3.211e-06 | 2.9275e-06 | 2.6816e-06 | 2.4693e-06 | 2.2859e-06 | 2.1263e-06 | 1.9851e-06 | 1.8566e-06 | 1.7344e-06 | 1.6116e-06 | 1.479e-06 | 1.3211e-06 | 1.0972e-06
R(surface), Ω | 0.0026088 | 0.0029041 | 0.0032305 | 0.0035899 | 0.0039838 | 0.0044126 | 0.0048752 | 0.0053689 | 0.0058889 | 0.0064288 | 0.0069817 | 0.007542 | 0.008108 | 0.0086845 | 0.0092858 | 0.0099398 | 0.010697 | 0.011656 | 0.01305 | 0.015713
Geom Factor | 3.353 | 4.1549 | 5.1413 | 6.3491 | 7.8189 | 9.5925 | 11.7094 | 14.2011 | 17.0851 | 20.3615 | 24.0141 | 28.023 | 32.3867 | 37.1562 | 42.4798 | 48.6742 | 56.3733 | 66.9348 | 83.9017 | 121.6349
Q | 1285.2586 | 1430.7155 | 1591.5028 | 1768.5891 | 1962.6521 | 2173.8918 | 2401.8153 | 2645.0418 | 2901.2218 | 3167.2048 | 3439.5772 | 3715.6001 | 3994.433 | 4278.4582 | 4574.7013 | 4896.8964 | 5269.9739 | 5742.4575 | 6429.2012 | 7741.0663
DOUBLE post with gap centered
r1 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025 | 0.025
r0/r1 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
r0 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015
h | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075
d | 1.9197e-05 | 2.9667e-05 | 4.5848e-05 | 7.0854e-05 | 0.0001095 | 0.00016922 | 0.00026151 | 0.00040413 | 0.00062455 | 0.00096518 | 0.0014916 | 0.0023051 | 0.0035623 | 0.0055052 | 0.0085077 | 0.013148 | 0.020319 | 0.0314 | 0.048526 | 0.074992
Cr, Farad | 2.2875e-12 | 2.1719e-12 | 2.0562e-12 | 1.9406e-12 | 1.825e-12 | 1.7094e-12 | 1.5937e-12 | 1.4781e-12 | 1.3625e-12 | 1.2469e-12 | 1.1313e-12 | 1.0156e-12 | 9.0002e-13 | 7.8439e-13 | 6.6877e-13 | 5.5315e-13 | 4.3753e-13 | 3.219e-13 | 2.0628e-13 | 9.0659e-14
Cg, Farad | 3.2602e-10 | 2.1096e-10 | 1.3651e-10 | 8.8332e-11 | 5.7158e-11 | 3.6986e-11 | 2.3933e-11 | 1.5487e-11 | 1.0021e-11 | 6.4845e-12 | 4.196e-12 | 2.7151e-12 | 1.7569e-12 | 1.1369e-12 | 7.3565e-13 | 4.7602e-13 | 3.0803e-13 | 1.9932e-13 | 1.2898e-13 | 8.3457e-14
L, Henry | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09 | 7.6624e-09
f0, MHz | 100.3457 | 124.5413 | 154.4584 | 191.3637 | 236.7414 | 292.2864 | 359.8659 | 441.4332 | 538.8881 | 653.8991 | 787.7479 | 941.3228 | 1115.4446 | 1311.7316 | 1534.2287 | 1792.2321 | 2105.713 | 2518.4127 | 3140.144 | 4357.3065
λ, meter | 2.9876 | 2.4072 | 1.9409 | 1.5666 | 1.2663 | 1.0257 | 0.83307 | 0.67913 | 0.55632 | 0.45847 | 0.38057 | 0.31848 | 0.26876 | 0.22855 | 0.1954 | 0.16727 | 0.14237 | 0.11904 | 0.095471 | 0.068802
Cu Skin δ, m | 6.5969e-06 | 5.9215e-06 | 5.3172e-06 | 4.777e-06 | 4.2949e-06 | 3.8653e-06 | 3.4835e-06 | 3.1452e-06 | 2.8467e-06 | 2.5842e-06 | 2.3545e-06 | 2.1539e-06 | 1.9786e-06 | 1.8246e-06 | 1.6871e-06 | 1.561e-06 | 1.4401e-06 | 1.3168e-06 | 1.1793e-06 | 1.0011e-06
R(surface), Ω | 0.0026133 | 0.0029114 | 0.0032423 | 0.0036089 | 0.0040141 | 0.0044602 | 0.004949 | 0.0054813 | 0.0060561 | 0.0066712 | 0.0073222 | 0.0080042 | 0.0087131 | 0.0094486 | 0.010219 | 0.011044 | 0.011971 | 0.013092 | 0.014619 | 0.017221
Geom Factor | 3.3646 | 4.1759 | 5.179 | 6.4165 | 7.938 | 9.8005 | 12.0664 | 14.8014 | 18.0691 | 21.9254 | 26.4134 | 31.5629 | 37.4012 | 43.9828 | 51.4432 | 60.0941 | 70.6052 | 84.4432 | 105.29 | 146.1019
Q | 1287.4783 | 1434.3239 | 1597.336 | 1777.953 | 1977.5503 | 2197.3275 | 2438.1523 | 2700.368 | 2983.5943 | 3286.5924 | 3607.3153 | 3943.2996 | 4292.5413 | 4654.928 | 5034.2587 | 5441.1112 | 5897.8005 | 6449.9142 | 7202.2015 | 8483.9781

## Optimize Rcav = .025
Optimize for maximum gap and maximum Q at 100 MHz, using goal seek on f = 100 with gap as variable, symmetric double-post re-entrant cavity with gap centered in middle
Fix cavity radius at .025 and vary both the post radius and cavity height
Post radius | 0.02 | 0.02 | 0.02 | 0.02 | 0.0175 | 0.0175 | 0.0175 | 0.0175 | 0.015 | 0.015 | 0.015 | 0.015 | 0.01 | 0.01 | 0.01 | 0.01 | 0.005 | 0.005 | 0.005 | 0.005 | 0.0025 | 0.0025 | 0.0025 | 0.0025
Inside height | 0.1 | 0.075 | 0.05 | 0.025 | 0.1 | 0.075 | 0.05 | 0.025 | 0.1 | 0.075 | 0.05 | 0.025 | 0.1 | 0.075 | 0.05 | 0.025 | 0.1 | 0.075 | 0.05 | 0.025 | 0.1 | 0.075 | 0.05 | 0.025
gap for 100 MHz, µm | 19.7 | 14.8 | 9.8 | 4.9 | 24.2 | 18.1 | 12 | 6 | 25.5 | 19.1 | 12.7 | 6.33 | 20.4 | 15.2 | 10.1 | 5.05 | 8.9 | 6.7 | 4.44 | 2.22 | 3.2 | 2.4 | 1.6 | 0.79
Q at 100 MHz | 715 | 704 | 683 | 626 | 1035 | 1012 | 969 | 859 | 1323 | 1285 | 1216 | 1048 | 1752 | 1686 | 1570 | 1300 | 1790 | 1722 | 1600 | 1321 | 1434 | 1390 | 1310 | 1116
f when gap = height | 4060 | 4381 | 4790 | 5316 | 4173 | 4467 | 4823 | 5222 | 4082 | 4355 | 4668 | 4978 | 3186 | 3460 | 3795 | 4193 | 2320 | 2580 | 2941 | 3502 | 2111 | 2366 | 2738 | 3371
Q when gap - height | 4556 | 4659 | 4725 | 4566 | 6687 | 6765 | 6729 | 6207 | 8452 | 8482 | 8311 | 7393 | 9886 | 9921 | 9671 | 8419 | 8620 | 8746 | 8678 | 7818 | 6987 | 6761 | 6853 | 6482
Fix cavity radius at .025 and cavity height at .075 m and vary the post radius
Best compromise between maximum gap & maximum Q at 100 MHz for post radius is about .010 m | [9 empty] | rPost/rCav= |  | 0.4
Post radius | 0.0025 | 0.005 | 0.006 | 0.007 | 0.0071 | 0.0072 | 0.0075 | 0.008 | 0.01 | 0.011 | 0.012 | 0.014 | 0.015 | 0.016 | 0.0175 | 0.02
gap for 100 MHz, µm | 2.39 | 6.68 | 8.54 | 10.4 | 10.5 | 10.7 | 11.3 | 12.1 | 15.2 | 16.5 | 17.6 | 18.9 | 19.1 | 18.9 | 18.1 | 14.8
Q at 100 MHz | 1390 | 1722 | 1765 | 1777 | 1777 | 1776 | 1774 | 1765 | 1686 | 1626 | 1554 | 1383 | 1285 | 1181 | 1012 | 704
f when gap = height | 5621 | 4702 | 4483 | 4376 | 4365 | 4345 | 4381 | 4274 | 4170 | 3398 | 4179 | 4267 | 4411 | 4489 | 4682 | 5381
Q when gap = height | 10420 | 11806 | 11817 | 11755 | 11739 | 11709 | 11740 | 11539 | 10501 | 9830 | 10049 | 9035 | 8536 | 7911 | 6926 | 5163
rPost/rCavity | 0.1 | 0.2 | 0.24 | 0.28 | 0.284 | 0.288 | 0.3 | 0.32 | 0.4 | 0.44 | 0.48 | 0.56 | 0.6 | 0.64 | 0.7 | 0.8
Fix cavity radius at .025 and cavity height at .01 m and vary the post radius
Best compromise between maximum gap & maximum Q at 100 MHz for post radius is about .010 m
Post radius | 0.0025 | 0.005 | 0.006 | 0.007 | 0.0071 | 0.0072 | 0.0075 | 0.008 | 0.01 | 0.011 | 0.012 | 0.014 | 0.015 | 0.016 | 0.0175 | 0.02
gap for 100 MHz, µm | 3.2 | 8.9 | 11.4 | 13.9 | 14.1 | 14.3 | 15 | 16.2 | 20.4 | 22.1 | 23.5 | 25.2 | 25.5 | 25.3 | 24.2 | 19.7
Q at 100 MHz | 1434 | 1790 | 1836 | 1849 | 1849 | 1849 | 1846 | 1836 | 1752 | 1687 | 1610 | 1426 | 1323 | 1212 | 1035 | 715
f when gap = height | 5324 | 4398 | 4222 | 4099 | 4089 | 4079 | 4052 | 4013 | 3924 | 3912 | 3921 | 4001 | 4075 | 4177 | 4396 | 5060
Q when gap = height | 10462 | 11867 | 11932 | 11840 | 11824 | 11807 | 11749 | 11633 | 10972 | 10549 | 10079 | 9022 | 8444 | 7834 | 6864 | 5085

## rCav .05
Re-entrant cylindrical RESULTS (no formulas). Single: one cylindrical post with gap at one end. Double: two posts of equal length at each end with gap in middle.
|  |  |  | Inside radius of cavity | radius of post | inside height of cavity | gap for 100 MHz | gap geo-metric spacer | full range chart x | full range chart left | full range chart right | low range chart x | low range chart left | low range chart right
|  |  |  | r2 | r1 | l | d | dmult
Input parameters |  |  |  | 0.05 | 0.03 | 0.15 | 0.00015877 | 1.4342 | 0.15 | 2400 | 1200 | 0.006 | 600 | 6000
Summary of Results | [6 empty] | (r1/r2) = | 0.6 | [11 empty] | max d
d | 0.00015877 | 0.0002277 | 0.00032656 | 0.00046834 | 0.00067167 | 0.00096327 | 0.0013815 | 0.0019812 | 0.0028414 | 0.004075 | 0.0058441 | 0.0083814 | 0.01202 | 0.017239 | 0.024723 | 0.035456 | 0.05085 | 0.072926 | 0.10459 | 0.14999
f0 single | 99.9997 | 118.7694 | 140.6871 | 166.0883 | 195.2566 | 228.3885 | 265.5626 | 306.7288 | 351.7342 | 400.3993 | 452.6509 | 508.7052 | 569.2999 | 635.9989 | 711.6638 | 801.3519 | 914.3435 | 1069.5402 | 1313.5878 | 1813.9336
f0 double | 101.1674 | 120.6382 | 143.6511 | 170.7378 | 202.4533 | 239.3498 | 281.9438 | 330.6798 | 385.9044 | 447.869 | 516.7888 | 592.985 | 677.1377 | 770.6874 | 876.4664 | 999.7953 | 1150.7144 | 1349.4465 | 1643.1352 | 2178.7798
Q single | 2570.5133 | 2801.3853 | 3048.9317 | 3312.7603 | 3591.893 | 3884.7013 | 4188.9381 | 4501.92 | 4820.896 | 5143.5994 | 5468.9275 | 5797.6709 | 6133.2554 | 6482.5917 | 6857.3753 | 7276.6607 | 7772.7583 | 8406.5738 | 9316.4408 | 10947.9089
Q double | 2585.478 | 2823.3384 | 3080.8814 | 3358.8096 | 3657.4877 | 3976.8295 | 4316.2016 | 4674.3834 | 5049.6392 | 5439.9634 | 5843.5535 | 6259.5385 | 6688.9659 | 7136.0783 | 7610.0605 | 8127.8566 | 8719.7545 | 9442.746 | 10419.7458 | 11998.5053
single λ , m | 2.9979 | 2.5242 | 2.1309 | 1.805 | 1.5354 | 1.3126 | 1.1289 | 0.97739 | 0.85233 | 0.74873 | 0.6623 | 0.58932 | 0.5266 | 0.47137 | 0.42126 | 0.37411 | 0.32788 | 0.2803 | 0.22822 | 0.16527
dble λ , m | 2.9633 | 2.4851 | 2.0869 | 1.7559 | 1.4808 | 1.2525 | 1.0633 | 0.90659 | 0.77686 | 0.66938 | 0.58011 | 0.50557 | 0.44273 | 0.38899 | 0.34205 | 0.29985 | 0.26053 | 0.22216 | 0.18245 | 0.1376
sngl MHz/µm | 0.2723 | 0.22171 | 0.17916 | 0.14345 | 0.11362 | 0.08889 | 0.068637 | 0.052322 | 0.03945 | 0.029535 | 0.022093 | 0.016653 | 0.012781 | 0.01011 | 0.008356 | 0.0073403 | 0.00703 | 0.0077082 | 0.011019
dbl MHz/µm | 0.28247 | 0.23279 | 0.19105 | 0.15598 | 0.12653 | 0.10185 | 0.081258 | 0.064203 | 0.050231 | 0.038956 | 0.030031 | 0.023127 | 0.017926 | 0.014134 | 0.01149 | 0.0098042 | 0.009002 | 0.0092761 | 0.011797
`
SINGLE post
r2 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
r1 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
r1/r2 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
l | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
d | 0.00015877 | 0.0002277 | 0.00032656 | 0.00046834 | 0.00067167 | 0.00096327 | 0.0013815 | 0.0019812 | 0.0028414 | 0.004075 | 0.0058441 | 0.0083814 | 0.01202 | 0.017239 | 0.024723 | 0.035456 | 0.05085 | 0.072926 | 0.10459 | 0.14999
α, deg | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90
α, rad | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708 | 1.5708
͞͞r0 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
r0 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
lM | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664 | 0.075664
C1, Farad | 7.6145e-12 | 7.2314e-12 | 6.8483e-12 | 6.4652e-12 | 6.0821e-12 | 5.699e-12 | 5.3159e-12 | 4.9327e-12 | 4.5496e-12 | 4.1665e-12 | 3.7834e-12 | 3.4003e-12 | 3.0172e-12 | 2.6341e-12 | 2.251e-12 | 1.8679e-12 | 1.4848e-12 | 1.1017e-12 | 7.1855e-13 | 3.3544e-13
C0, Farad | 1.5768e-10 | 1.0994e-10 | 7.6662e-11 | 5.3454e-11 | 3.7272e-11 | 2.5989e-11 | 1.8122e-11 | 1.2636e-11 | 8.8107e-12 | 6.1435e-12 | 4.2837e-12 | 2.9869e-12 | 2.0827e-12 | 1.4522e-12 | 1.0126e-12 | 7.0607e-13 | 4.9233e-13 | 3.4329e-13 | 2.3937e-13 | 1.6691e-13
L, Henry | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08
f0, MHz | 99.9997 | 118.7694 | 140.6871 | 166.0883 | 195.2566 | 228.3885 | 265.5626 | 306.7288 | 351.7342 | 400.3993 | 452.6509 | 508.7052 | 569.2999 | 635.9989 | 711.6638 | 801.3519 | 914.3435 | 1069.5402 | 1313.5878 | 1813.9336
λ, meter | 2.9979 | 2.5242 | 2.1309 | 1.805 | 1.5354 | 1.3126 | 1.1289 | 0.97739 | 0.85233 | 0.74873 | 0.6623 | 0.58932 | 0.5266 | 0.47137 | 0.42126 | 0.37411 | 0.32788 | 0.2803 | 0.22822 | 0.16527
Cu Skin δ, m | 6.6083e-06 | 6.0637e-06 | 5.5713e-06 | 5.1276e-06 | 4.7292e-06 | 4.3727e-06 | 4.0551e-06 | 3.7732e-06 | 3.5235e-06 | 3.3025e-06 | 3.106e-06 | 2.9299e-06 | 2.7696e-06 | 2.6203e-06 | 2.4771e-06 | 2.3344e-06 | 2.1854e-06 | 2.0206e-06 | 1.8233e-06 | 1.5516e-06
R(surface), Ω | 0.0026088 | 0.0028431 | 0.0030944 | 0.0033621 | 0.0036454 | 0.0039426 | 0.0042514 | 0.004569 | 0.0048928 | 0.0052203 | 0.0055505 | 0.0058841 | 0.0062247 | 0.0065792 | 0.0069596 | 0.0073851 | 0.0078886 | 0.0085319 | 0.0094553 | 0.011111
Geom Factor | 6.706 | 7.9648 | 9.4346 | 11.138 | 13.094 | 15.3159 | 17.8088 | 20.5694 | 23.5875 | 26.851 | 30.3551 | 34.1141 | 38.1776 | 42.6505 | 47.7246 | 53.7392 | 61.3164 | 71.724 | 88.09 | 121.6435
Q | 2570.5133 | 2801.3853 | 3048.9317 | 3312.7603 | 3591.893 | 3884.7013 | 4188.9381 | 4501.92 | 4820.896 | 5143.5994 | 5468.9275 | 5797.6709 | 6133.2554 | 6482.5917 | 6857.3753 | 7276.6607 | 7772.7583 | 8406.5738 | 9316.4408 | 10947.9089
DOUBLE post with gap centered
r1 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
r0/r1 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
r0 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
h | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
d | 0.00015877 | 0.0002277 | 0.00032656 | 0.00046834 | 0.00067167 | 0.00096327 | 0.0013815 | 0.0019812 | 0.0028414 | 0.004075 | 0.0058441 | 0.0083814 | 0.01202 | 0.017239 | 0.024723 | 0.035456 | 0.05085 | 0.072926 | 0.10459 | 0.14999
Cr, Farad | 3.8208e-12 | 3.6293e-12 | 3.4377e-12 | 3.2462e-12 | 3.0546e-12 | 2.8631e-12 | 2.6715e-12 | 2.4799e-12 | 2.2884e-12 | 2.0968e-12 | 1.9053e-12 | 1.7137e-12 | 1.5222e-12 | 1.3306e-12 | 1.1391e-12 | 9.4751e-13 | 7.5595e-13 | 5.644e-13 | 3.7284e-13 | 1.8129e-13
Cg, Farad | 1.5768e-10 | 1.0994e-10 | 7.6662e-11 | 5.3454e-11 | 3.7272e-11 | 2.5989e-11 | 1.8122e-11 | 1.2636e-11 | 8.8107e-12 | 6.1435e-12 | 4.2837e-12 | 2.9869e-12 | 2.0827e-12 | 1.4522e-12 | 1.0126e-12 | 7.0607e-13 | 4.9233e-13 | 3.4329e-13 | 2.3937e-13 | 1.6691e-13
L, Henry | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08 | 1.5325e-08
f0, MHz | 101.1674 | 120.6382 | 143.6511 | 170.7378 | 202.4533 | 239.3498 | 281.9438 | 330.6798 | 385.9044 | 447.869 | 516.7888 | 592.985 | 677.1377 | 770.6874 | 876.4664 | 999.7953 | 1150.7144 | 1349.4465 | 1643.1352 | 2178.7798
λ, meter | 2.9633 | 2.4851 | 2.0869 | 1.7559 | 1.4808 | 1.2525 | 1.0633 | 0.90659 | 0.77686 | 0.66938 | 0.58011 | 0.50557 | 0.44273 | 0.38899 | 0.34205 | 0.29985 | 0.26053 | 0.22216 | 0.18245 | 0.1376
Cu Skin δ, m | 6.57e-06 | 6.0165e-06 | 5.5136e-06 | 5.0573e-06 | 4.6444e-06 | 4.2714e-06 | 3.9356e-06 | 3.634e-06 | 3.3639e-06 | 3.1226e-06 | 2.9069e-06 | 2.7137e-06 | 2.5395e-06 | 2.3804e-06 | 2.2321e-06 | 2.0899e-06 | 1.9481e-06 | 1.7989e-06 | 1.6302e-06 | 1.4157e-06
R(surface), Ω | 0.002624 | 0.0028654 | 0.0031268 | 0.0034089 | 0.003712 | 0.0040361 | 0.0043805 | 0.0047441 | 0.0051249 | 0.0055211 | 0.0059307 | 0.0063529 | 0.0067887 | 0.0072425 | 0.0077235 | 0.008249 | 0.0088497 | 0.0095835 | 0.010575 | 0.012177
Geom Factor | 6.7844 | 8.0901 | 9.6333 | 11.4498 | 13.5766 | 16.0509 | 18.9073 | 22.1756 | 25.879 | 30.0344 | 34.6562 | 39.7659 | 45.4093 | 51.6828 | 58.7764 | 67.0469 | 77.1676 | 90.4947 | 110.1897 | 146.1103
Q | 2585.478 | 2823.3384 | 3080.8814 | 3358.8096 | 3657.4877 | 3976.8295 | 4316.2016 | 4674.3834 | 5049.6392 | 5439.9634 | 5843.5535 | 6259.5385 | 6688.9659 | 7136.0783 | 7610.0605 | 8127.8566 | 8719.7545 | 9442.746 | 10419.7458 | 11998.5053

## Copper tubing sizes
Standard Copper Tube sizes
https://en.wikipedia.org/wiki/Copper_tubing#United_States,_Canada,_and_Brazil
` | Outside diameter | Inside diameter (ID)
size | (OD) | [in (mm)]
| [in (mm)] | Type K | Type L | Type M
1⁄4 | 3⁄8 (9.5) | 0.305 (7.747) | 0.315 (8.001)
3⁄8 | 1⁄2 (12.7) | 0.402 (10.211) | 0.430 (10.922) | 0.450 (11.430)
1⁄2 | 5⁄8 (15.875) | 0.528 (13.411) | 0.545 (13.843) | 0.569 (14.453)
5⁄8 | 3⁄4 (19.05) | 0.652 (16.561) | 0.668 (16.967) | 0.690 (17.526)
3⁄4 | 7⁄8 (22.225) | 0.745 (18.923) | 0.785 (19.939) | 0.811 (20.599)
1 | 1+1⁄8 (28.575) | 0.995 (25.273) | 1.025 (26.035) | 1.055 (26.797)
11⁄4 | 1+3⁄8 (34.925) | 1.245 (31.623) | 1.265 (32.131) | 1.291 (32.791)
11⁄2 | 1+5⁄8 (41.275) | 1.481 (37.617) | 1.505 (38.227) | 1.527 (38.786)
2 | 2+1⁄8 (53.975) | 1.959 (49.759) | 1.985 (50.419) | 2.009 (51.029)
21⁄2 | 2+5⁄8 (66.675) | 2.435 (61.849) | 2.465 (62.611) | 2.495 (63.373)
3 | 3+1⁄8 (79.375) | 2.907 (73.838) | 2.945 (74.803) | 2.981 (75.717)